1. Maximum Sample Rate Test

Verify the highest speed at which the hardware can reliably capture data without dropping bits.

- Method: Input a high-frequency square wave (e.g., from a signal generator).

- Goal: Ensure the captured waveform matches the input frequency without aliasing.

- Check: Verify the Shannon-Nyquist limit (sample at least 2x the input frequency). [1]

2. Pulse Width & Glitch Capture

Test the minimum pulse width the analyzer can detect. [1]

- Method: Input a series of narrowing pulses (100ns, 50ns, 10ns, etc.).

- Goal: Identify the "blind spot" where the hardware fails to register a transition.

- Check: Confirm if the hardware can capture pulses shorter than the sample period. [1]

3. Multi-Channel Skew Test

Check for timing alignment issues between different input pins.

- Method: Send the same fast clock signal to all input channels simultaneously.

- Goal: Measure the time delta between the first and last channel registering the edge.

- Check: Ensure skew is within your design's tolerance (ideally <1 sample period).

4. Trigger Latency & Reliability

Ensure the analyzer starts recording exactly when the trigger condition is met. [1]

- Method: Set a complex trigger (e.g., a specific I2C address or pattern) and measure the offset.

- Goal: Confirm the "trigger point" in the buffer aligns with the actual event.

- Check: Test "Pre-trigger" and "Post-trigger" buffer depth accuracy. [1, 2]

5. Threshold & Logic Level Stress

Validate the comparator or input buffer performance across voltage ranges.

- Method: Use a variable voltage source to find the exact high/low transition point.

- Goal: Verify the switching threshold matches your design (e.g., 1.8V, 3.3V, or 5V logic).

- Check: Test for "double-triggering" or noise on slow-rising edges (hysteresis check). [1]

6. Sustained Data Throughput

Check for buffer overflows or USB/Interface bottlenecks.

- Method: Capture a high-speed data stream (like an SPI flash read) for several seconds.

- Goal: Ensure no data gaps occur during long-duration captures.

- Check: Monitor for "Buffer Overflow" errors in your software or FPGA logic.

- Maximum Sample Rate Test

- Threshold and Logic Trigger

- Minimum Pulse Width